## Slide 1
検出器
Yoshie OTAKE Riken center for advanced Photonics(RAP), Neutron beam team technology Team RIKEN
M.Takamura, Y.Yamagata, H.Sunaga, A.Taketani, Y.Ikeda, M.Yamada Y.Seki
H.Suzuki(JAEA), Y.Kumagai(Tokyo city Univ.) et al
RIKEN compact neutron source RANS and its applications
-RIKEN Accelerator-driven compact Neutron Source
8th June 2015
Jagiellonian Symposium
Krakow
Great thanks

## Slide 2: RANS project goals：CANS for practical use
Compact neutron source system easy and handy to use –floor-standing type
  industrial use,
    non-destructive inspection
    industrial material development analysis,
      Together with VCAD (simulation)
Fast neutron large area imaging system
X-ray
Cooling pomp
Power suppl
High frequency gene.
Transportable X-ray instrument for bridges Prof.M.UESAKA (Univ.Tokyo) 3.95MeV
X-band e-linac
      Transportable non-destructive inspection of large scale bridges
RANS-RIKEN Accelerator-driven compact Neutron Source

## Slide 3: RANS=RIKEN Accelerator-driven compact Neutron Source development policy : easy to access on-site
=low energy Be(p,n) reaction
<Techniques overcome and features>・・・・・・・・・
1.Small system including shielding-> proton linac
Easy maintenance
  Neutron flux ~1012/sec@７MeV, 100μA(max)
  Is nearly 30 times more than with 3.5MeV (proton nergy)
2011started. 2012 construction Main members
Space ~4m*15m
  ７MeV、100μA、Rf power supply.: 350kW(peak) duty 1.3%,Electric power peak 40kVA,
  Cooling water：75L/min ,pulse width (30～200μs)repetition frequency～20～180Hz
  RF power 425MHz, Injection energy0.030-3.5MeV
  2.Long Life Be target invention by Dr.Y.Yamagata
3.compact and low cost shielding design
  Multilayer shielding of target station
★blistering (hydrogen embrittlement damage)
=low energy Be(p,n) reaction
7MeV proton

## Slide 4
Proton Linac’7MeV
Q DTL RFQ
Neutron Beamline
：Be(Target)
：V(backing) Hydrogen diffusion
：PE(polyethylene)Moderator
：C(graphite) Reflector
：BPE(Boron PE )Neutron Shielding
：Pｂ γ-ray shielding
②
④
⑤
⑥
③
①
Inside Target station
Neutron detector,
sample box
Target station
RANS compact Neutron source for realization of practical use, and development of trans-portable instrument for infrastructure safety
Neutron
Be
Proton 7MeV
2013 January,
RANS starts running
After ten month construction.
Reflector
Graphite
BPE
Be
Neutron
Proton
Moderator
absorber

## Slide 5: RANS Neutron spectrum (PHITS simulation)
@ maximum averaged current 100μA
Lower peak energy
E~50meV λ~0.13nm
thermal neutron
E=25meV, λ=1.8Å
higher peak energy
E~1.5MeV
fast neutron
Development for the fast neutron imaging detector
：Be(Target)
：V(backing) Hydrogen diffusion
：PE(polyethylene)Moderator
：C(graphite) Reflector
：BPE(Boron PE )Neutron Shielding
：Pb γ-ray shielding
Reflector
Graphite
Moderator
absorber
BPE
Proton
Be
N

## Slide 6
Be(p,n) reaction seems to be ideal for low energy neutron source, while major problem in Be target is the blistering (hydrogen embrittlement damage) of target caused by low energy proton beam.
We proposed a new structure target with backing of highly hydrogen diffusible material.
The performance estimation and design was done by using ion injection simulation (SRIM), and finite element simulation (hydrogen diffusion, thermo-fluid analysis, structural analysis).
We expect that the newly-designed target has life time more than several thousand hours.
Target destruction by blistering
Long life Be(p,n) target development Yutaka Yamagata RIKEN
P
Ｎ
7 MeV proton
Be
B
Neutron
Y.Yamagata
↓V backing +Be
Patent PCT/HP2013/056188

## Slide 7
Long life Be(p,n) target Yutaka Yamagata RIKEN
Y.Yamagata, K.Hirota, J.Ju, S.Wang, S.Morita, J.Kato, Y.Otake, A.Taketani, Y.Seki, Y.Yamada, H.Ota, U.Bautista, Q.Jia
“Development of a neutron generating target for compact neutron soruces using low energy proton beams” Journal of Radioanalytical and Nuclear Chemistry
DOI 10.1007/s10967-015-4059-8
Picture taken in Feb. of 2014
after about 3900μAh
Neutron
Proton 7MeV
After 2 years, ~9100 μA H
NO cracks or blistering is observed

## Slide 8: A example of industrial use,
Steel
Compact neutron source is useful!!
Compact neutron Application results for practical use.

## Slide 9
Alloy
Steel
蒸留水中浸漬処理110分
エアブロー乾燥処理30分
(位置固定）
中性子
イメージング
測定 (10分）
中性子
イメージング
測定 (10分）
中性子
イメージング
測定 (10分）
KOBELCO
Dr. T . Nakayama
A.Taketani H.Sunaga, M.Yamada, S.Mihara
RANS take on a challenge,
Non-destructive observation of inside the steel
under the film
RANS on-site
ISIJ (the Ion and Steel Institute of Japan) Research Activity,
Nippon Steel and Sumitomo Metal
JFE-Steel
Kobelco
Daido Steel
Univ. Inst. (Hokkaido, Tohoku, Ibaraki, Kyoto, Tokyo city, Toyohashi, Tokyo metropolitan, CROSS)
Neutron imaging detector
Scintillator + C CCD

## Slide 10
Corrosion Visualization
RANS imaging time~5~20 min.
Alloy
Steel
Whether it is possible to visualize the steel corrosion under-film?
base material is Ｆｅ
What we should find is ＦｅＯＯＨ, Fe2O3, Fe3O4
Film contains a lot of H,
→It is impossible for X-ray to distinguish the corrosion from base material Fe
Using neutrons can we distinguish corrosion, rust under the film, difference between steel and alloy?
KOBELCO
Proposed by Dr. T . Nakayama
Industrial application- Ion and steel

## Slide 11
110minutes wet
Success of the visualization wet-dry process
Alloy
Steel
Normal condition( in the room)
30minutes trying

## Slide 12
The ratio of projection into horizontal axis of neutron intensity taken after water immersion treatment and after drying treatment. Only artificial defect parts of E16 (open circle) and S18 (closed circle) were projected, respectively.
RANS Success of non-destructive visualization of the rust in the steel under-film different kinds of steels and moisture reduction

## Slide 13
3D information of the corrosion in the steel under-film
Success of moisture reduction for long time
alloy
Steel
wet
1 hour
2 hours
Corrosion with water
Dr.Taketani
RANS explosure time~5 min per one shot.
Low energy neutrons of compact neutron source is powerful tool!!

## Slide 14: Metallographic observation, plastic forming process on 13th Y.Ikeda has presented
Neutron diffraction
with compact neutron source??
Pulsed neutron TOF measurement
as received
After compression
Electron Back Scattering Diffraction Only surface information Pure Titanium (together with Prof. T.Hama Kyoto Univ)

## Slide 16
RIKEN has the special Simulation technique : Body Side Outer Sheet metal forming simulation the largest parts
Body Side Outer
Plastic deformation under press process, to prevent racks, wrinkles, FEM analysis is useful tool.
VCAD project RIKEN 2001-2011
Makinouchi M.Takamura, H.Sunaga
bulk metal observation on site should be possible only with neutron beam

## Slide 17
Cut
30 plates with the same direction
~10
~10
>30
Load direction of elongation
引張試験片
W=30mm、t=1.2mm
Sample with 19.4% elongation
JSC440W
Nominal stress
First trial experiment at RANS was done 30 31 July 2014
Stress-strain curve
Nominal strain

## Slide 18
Neutrons
Diffraction results by RANS
タイムフォーカシング
負荷方向の回折を測定
JSC440W
10 minutes measurement
Before elongation
Ditector
Neutron
Cut
W=30mm、t=1.2mm
JSC440W
BCC steel sample
JSC440W
~10
~10
>30
JAEA H.Suzuki,
RIKEN Takamura, Y.IKEDA

## Slide 19
入射エネルギースペクトルで規格化
タイムフォーカシング
JSC440W
負荷方向の回折を測定
90分測定→２０分から１０分で様子見える
As-receive（0%）rather strong orientation along (110)
Under elongation along 110 orientation
Success of the observation of texture @ RANS
Stress-strain curve
Lattice constant [nm]
JAEA H. Suzuki,
RIKEN Takamura, Y.IKEDA
Neutron
success of the observation of texture component
After 19% elongation
Neutrons

## Slide 21: Neutron diffraction → texture component (10min.measurement)
→Whether the polar figure can be obtained with compact neutron source with reasonable time?
(110)
steel
Neutron beam

## Slide 22
Neutron beam
Crystal texture information of all direction for one crystal orientation
Change of crystal texture structures before and after the plastic deformation can be observed.
Compact neutron source will be new tool for industrial use on-site!

## Slide 23
Fast neutron imaging system for nondestructive inspection of large-scale concrete structures
Y. Seki1，A. Taketani1, T. Hashiguchi1, H. Ota1, S. Tanaka2, K. Kino3, K. Hirota4, H. Baba1, S. Wang1,4, Y. Yamagata1, Y. Otake1
RIKEN1, KEK2, Hokkaido Univ.3, Nagoya Univ.4 , Xi’an Jiaotong Univ.5
1
Notes: Thank you, Mr. Chairman.
I will be talking about
"Fast neutron imaging system for nondestructive inspection of large-scale concrete structures".

## Slide 24
Background
Aging deterioration of large-scale concrete structures
  Lifespan of concrete  60 years  peak in 2025 in Japan  Lifetime expiration 42,000 bridges
  New construction of bridges/highways is impossible  Diagnosis, preventive maintenance, life extension
Collapse of the Ynys-y-Gwas brg. (UK, 1985)
Assessment of concrete deterioration
Deteriorated concrete
Width of Steel bar
Void
Water
Fracture of steel bar
Required resolution  3 cm
Fast neutron transmission imaging
as on-site nondestructive inspection technique
High penetration power
Sensitivity to water
3
Notes: So,
we have to diagnose concrete deterioration as soon as possible,
and extend its lifetime with preventive maintenance.
The concrete deterioration is evaluated
by observing inside concrete,
observing cracks, water,
rust on steel and fractures of steel.
Fracture of steel can directly cause the collapse of the structures.
For this nondestructive inspection of concrete,
Civil engineering researchers require
the spatial resolution of 3 cm.
For such purpose,
we propose a new method of the fast neutron imaging.
The fast neutron, with the energy of a few MeV, is a promising probe
because it can penetrate thick concrete
and is sensitive to water.

## Slide 25: Transportable compact neutron system
Steel bridges
Pre-stressed bridge,
PC bridge
steel brake
steel bridge
corrosion
External aspects
Internal aspect
we cannot know without pealing the concrete slab
Non-destructive inspection, use fast neutron beam (~MeV)
After disassemble joint part・・・

## Slide 26: Fast neutron (some MeV)：concrete slab with~50cm thickness
Slow neutrons)
E Energy |  | 106eV | 103eV | 25meV |  | 1meV | 0.01meV | 10-7eV
Temperature |  | 1010K |  | 300K |  | 10K | 0.1K | 10-5K
Velocity |  | 107m/s |  | 2200m/s |  |  | 1m/s | 
Wave length |  |  |  | 0.18nm |  | 1nm |  | 1μm
10-12m
Fast neutrons
E~1MeV,
Low absorption for almost all elements,
But still high scattering cross section for H atom
↓
Detection of Water in the concrete (~50cm)
Scattering cross section
Absorption cross section
Hydrogen and neutron nuclear reaction
Slow neutrons
(25 meV)
Fast neutrons *MeV
106

## Slide 27
Fast neutron detector for concrete inspection
Scintillators
Scintillators + MPPCs
Readout circuit
Neutron beam
Multi-Pixel Photon Counter (MPPC)
Plastic scintillator
HAMAMATSU S10931-050P
Effective area: 3 mm x 3 mm Counting ability: 3 mV/p.e. Time resolution: 500-600 ps
BC-408 3 cm(H) x 3 cm(W) x 5 cm(D)
Detector design
  Outdoor use
  Mobile use
  Mass production
2nd Prototype 4 x 4 ch (12 cm x 12 cm) detector
Scintillators + MPPCs
Plastic Scintillator
MPPC
Amplifier
Neutron
Comparator
Counter
PC
Accelerator T0
Network
5
Notes: Taking into account outdoor use, mobile use, and capability of mass production,
we designed the detector as simple as possible.
Our detector consists of plastic scintillators and
solid state Multi-Pixel Photon Counters (MPPCs).
The incident neutrons recoil protons in the scintillator,
and the recoiled protons produce scintillation light,
and eventually the MPPCs detects the light.
The cross-section of one scintillator is 3 cm high, 3 cm wide.
A prototype has 4 x 4 array structure,
thus, the total detector area is 12 cm x 12 cm.

## Slide 28
Success of observation difference of steel bar in the concrete
Insert ion bars into concrete
0, 1, 2, 3
Concrete 30cm
Concrete +Pb 10cm
Transmission pro.
・comparison with the experimental results and simulation by GEANT4.
Number of steel bar

## Slide 29
Fast neutron imaging through 30cm concrete block
(>1MeV)Non-destructive inspection
Proton energy: 7 MeV Beam current: 11 mA (avg.) Rep. rate: 20 Hz Pulse width: 100 ms
Proton linac
中性子線
Neutron
Neutron
30cm
30cm
Observation of the hall inside
Observation of the steel bar
Neutron
4*4pixel detector

## Slide 30
Steel observation
1
5
9
13
14
10
6
2
3
7
11
15
4
8
12
16
透過率
Position [mm]
8

## Slide 31: Towards the health diagnostics for the infrastructure

## Slide 32: Based on the imaging data what we can predict?
Φ12㎜
100mm
200mm
520mm
φ12㎜
φ16㎜
φ16㎜
Cement, concrete
（出典： www.hirosima.co.jp）
（出典： Google Map)
Example wooden house basement
VCAD project RIKEN
2001-2011
A. Makinouchi director

## Slide 33
Rotation and upward movement
X-ray
CT measurement parameter,
Upward pitch：1mm
×：520times
M time： 38hours
X-ray CT
a
b
c
d
e
f
Modeling of 3D data
a
b
c
d
e
f
空洞（巣）
Steel
粗骨材
Using industrial X-ray
CT instrument
VCAD project
2001-2011
Once we have 3D data,

## Slide 34
Outside shape
Inside elements
coarse aggregate inforced steel air hole
Extraction part of
Cement
Results from x-ray CT-> 3D shape modeling
（wooden house basement）
VCAD project
2001-2011
After these procedure,
Billion air holes become visible

## Slide 35: Analytical model tetrahedron elements method and give some constraints-> structural analysis simulation
tetrahedron elements model
Number of elements： 9,880,531
Number of contacts： 1,772,314
coarse aggregate
Air holes）
Push １２ｋN
Steel bar
Completely fixed
VCAD project2001-2011

## Slide 36: Analytical results （wooden house basement）
0.00
0.75
1.5
2.25
3.00
elastic deformation
analysis result
（distribution of stress）
MPa
X-ray CT data
Steel bar
VCAD project
2001-2011
We are able to predict, from where the crack will start with which stress, now.
From the measurement to prediction,
health diagnostics for infrastructure!
Without drawing for design (or different from design),
whenever we get such inside data as structure, picture, imaging, radiography, we can predict when or how it will destroy

## Slide 37: Trans portable compact neutron system in a truck will be realized in some years.
放射線障害等防止法第１０条 および 関連規定（平成１７年７月改定）Japanese regulation 4MeV>linac
  橋梁等の非破壊検査に用いる直線加速器で４メガ電子ボルト以上のエネルギーを有する放射線を発生しないものは、放射線発生装置の使用の場所の変更を都度許可を得る必要がなく届出で足りることとする。（ただし、設備については、事前に原子力規制委員会原子力規制庁の届け出許可が必要。）
The radiation level 3m far from the center of the target will be 0.6μSv/h by GEANT4 simulation.
N
50cm
50cm

## Slide 38: Collaboration and the support from JCANS (Prof. Kiyanagi), and
M.Furusaka
M.Onuma
H.M.Shimizu
Y.Iwashita,
N.Hayashizaki
T.Sato
W.Yashiro
Y.Tomota
H.Suzuki
S.Sato
S.Tasaki
K.Hirota,
Y.Shiota,
T.Nagae,
Hirose,
A.Momose,
T.Shinohara,
So many people
Steel companies,
RIKEN
Y.Yamagata,
J.Kato,
A.Taketani,
H.Sunaga,
M.Takamura,
S.Wang,
T.Hashiguchi,
Y.Seki,
M.Yamada,
M.Ikeda,
S.Mihara
H.Ohta,
S.Yanagimachi,
Many people
Dr.H.Suzuki, JAEA
Dr.Y.Kumagai,
Tokyo City Univ.
A.Prof.N.Hayashizaki
TITECH
Dr.T.Sakai
JAEA
Dr.N.Hagura,
Tokyo City Univ.

## Slide 39: New: neutron application facilities building in RIKEN

## Slide 40
①
②
Portable system
Non-destructive diagnostic system
Summary RANS, towards practical use
Compact neutron source for industrial use floor-standing type
Development for the transportable neutron system
Fast neutron imaging detector,
health diagnosis system development
3D neutron imaging (3D)
Neutron diffraction for metal deformation
Large are detector
Crystal texture,
pole figure measurement

## Slide 42
Success of the observation of steels, void
New fast neutron imaging detector development
1pixel = 30 mm x 30 mm 10 mm x 10 mm steel Φ18 mm void (air hole)
1m*1m 1024 channner large area fast neutron imaging detector is now being tested with real bridge
1平米
橋梁用1024チャンネル検出器
[32 x 8 ピクセル (1ピクセル = 20 mm x 20 mm)]