## Slide 1: Versatile Compton camera for high energy gamma rays: Monte Carlo comparison with Anger camera for medical imaging
Mattia Fontana
2017 – 06 - 08
2nd jagiellonian symposium - krakow

## Slide 2: OUTLINE
  INTRODUCTION
    Nuclear medicine: PET & SPECT
    Radionuclides
    Compton camera principle
    Anger vs Compton for SPECT
  COMPTON CAMERA DEVELOPMENT
    CLaRyS prototype
  SIMULATION STUDY
    Comparison between Compton and Anger cameras for medical imaging

## Slide 3: INTRODUCTION

## Slide 4
Nuclear medicine
Diagnosis/therapy techniques based on radioemitters injected into the patient body (bloodstream, inhaled, swallowed)
Diagnosis: physiological information based on SPECT (Single Photon Emission Computed Tomography) or PET (Positron Emission Tomography) imaging
SPECT
Single photon detection from gamma emitters
Clinical routine based on Anger cameras (Hal Anger 1958)
Mechanical collimator coupled to scintillating-segmented photodetector
  low detection efficiency
  limited energy range exploitable
    low energy sources
PET
b+ annihilation: 2 back-to-back 511 keV gammas
Clinical system based on a ring of scintillating-segmented photodetectors
Coincidence detection: no need for mechanical collimation
  good detection efficiency
Positron diffusion before annihilation:
  spatial accuracy affected

## Slide 5
Higher energy sources [Nurdan 2015]:
24Na, 42K, 59Fe, 47Ca, 65Zn, 28Mg, 212Bi, 208Tl, 132I, 91mY, ...
Vectorised targeted therapy:
Beta/alfa sources used for targeted cancer therapy often emit also high energy gamma rays (ex: 89Sr, 111At, 225Ac, …) with limited BR
  Pre-/per- treatment imaging for therapy monitoring
Radionuclides
Application | Isotope | Decay mode | g energy [keV] | Half-life
SPECT | Thallium 201 | g | 135 | 73 h
 | Technetium 99m | g | 140 | 6 h
 | Iodine 123 | g | 159 | 13 h
 | Iodine 131 | g | 364 | 8 d
PET | Fluorine 18 | b+ | 511 | 110 m
 | Oxygen 15 | b+ | 511 | 2 m
Common radioemitters for clinical use
What’s needed:
Handle higher energy photons
Higher efficiency
But…

## Slide 6
COMPTON CAMERAS
Common radioemitters for clinical use
Higher energy sources [Nurdan 2015]:
24Na, 42K, 59Fe, 47Ca, 65Zn, 28Mg, 212Bi, 208Tl, 132I, 91mY, ...
Vectorised targeted therapy:
Beta/alfa sources used for targeted cancer therapy often emit also high energy gamma rays (ex: 89Sr, 111At, 225Ac, …) with limited BR
  Pre-/per- treatment imaging for therapy monitoring
But…
Radionuclides
Application | Isotope | Decay mode | g energy [keV] | Half-life
SPECT | Thallium 201 | g | 135 | 73 h
 | Technetium 99m | g | 140 | 6 h
 | Iodine 123 | g | 159 | 13 h
 | Iodine 131 | g | 364 | 8 d
PET | Fluorine 18 | b+ | 511 | 110 m
 | Oxygen 15 | b+ | 511 | 2 m

## Slide 7
Compton camera: the principle
Two detector components:
Scatterer – low Z position sensitive detector
Absorber – high Z segmented detector
Two positions and energy deposits
Compton scattering formula
Reconstruction of a cone surface for each gamma event
Cones overlapping to retrieve the radioactive source distribution
E1, r1
E2, r2
E0
Seo 2009
Kim 2013

## Slide 8
g
Source
Collimator
Photodetector
ANGER vs COMPTON for SPECT
Scatterer
Absorber
Source
g
g
 | ANGER CAMERA | COMPTON CAMERA
Events | Single interaction | Time coincidences
Collimation | Physical | “Electronic”
Efficiency | Low | Intrinsically higher
Source energy | <= 364 keV | 100 keV – some MeV
Image reconstruction | Tomography (not needed for 2D) | Iterative algorithms (complex and long)
Image | 3D with tomography | Potentially 3D with single detector
Cost | Relatively low | High
Status | Commercially available | Several prototypes under test
[McKisson 1994]

## Slide 9: COMPTON CAMERA DEVELOPMENT

## Slide 10
Compton camera: the prototype
COMPTON CAMERA (CC)
Components: [Krimmer 2015]
silicon scatterer stack
  7 planes of DSSDs (Double-sided Silicon Strip Detectors)
  9x9x0.2 cm3 each
segmented BGO (Bismuth Germanate) absorber
  96 blocks
  3.5x3.8x3.0 cm3 each
  4 tube PMs per block
LM-MLEM (List-Mode MLEM) algorithm developed by the CREATIS group (Lyon)
Prototype design optimized by Monte Carlo simulations for ion beam therapy monitoring
(Roellinghoff 2011)
5 French institution involved in the development –
CLaRyS collaboration -> IPNL, LPSC, LPC, CPPM, CREATIS

## Slide 11: SIMULATION STUDY
Study objective and description

## Slide 12
CC
STATE OF THE ART
Comparison with commercial Anger camera performed at single gamma energy (364 keV) – Han 2008
  Encouraging results:
    Efficiency gain of a factor > 20
    Comparable spatial resolution for the same number of events
OBJECTIVE OF THE PRESENT WORK
Estimate of Compton camera performance in SPECT as a function of:
  Scatterer energy resolution and material
    Main development challenge
  Source activity and random coincidences
Benchmark Anger camera (AC) and Compton camera (CC):
  Comparison of main performance parameters in wide
  energy range (245 keV – 2614 keV)
Simulation study
AC

## Slide 13
GATE simulation of Infinia by GE Healthcare
Collimator HEGP (High Energy General Purpose)
  19x28x6.6 cm3
  Lead
  Hexagonal holes 2.0 mm radius
  1.8 mm septal thickness
NaI block
  19x28x1 cm3
  4 mm FWHM spatial resolution
  10% FWHM energy resolution @ 140 keV
  80 keV energy threshold
Back compartment
  19x28x2.5 cm3
  Glass
[HealthCare 2006]
19 cm
28 cm
10 cm
10 cm
21 cm
28 cm
6 cm
15 cm
Geant4 simulation of the CLaRyS Compton camera
Scatterer
  7 Si planes (500 e- ENC)
  1 cm distance between each plane
  FWHM resolutions: 20 ns time,~ 5 keV energy, 1 mm spatial
BGO Absorber
  21x28x3.0 cm3 (adapted size for efficiency comparison) – 8 x 6 block matrix
  FWHM resolutions: 3 ns time, 21% energy @ 667 keV, 4.4 mm spatial transverse plane
MLEM reconstruction method (CREATIS)
Point-like radioactive source study at different energies
Simulation settings

## Slide 14
Compton camera study for SPECT
SCATTERER ENERGY RESOLUTION
= 1 keV in Han 2008
2 energy resolution values tested
  = 2 keV close to design expectations
  = 4 keV test with preliminary electronics card
Comparison with alternative material – cadmium telluride (CdTe)
SOURCE ACTIVITY
True and random coincidences studied as a function of the source activity at fixed energy (555 keV)
Activity fixed to value of Anger-SPECT clinical interest
  200 MBq – 50 % of random coincidences
Random coincidence ratio studied as a function of primary energy at fixed activity
= 2 keV fixed for comparison study
Activity of 200 MBq for comparison study

## Slide 15
CC
STATE OF THE ART
Comparison with commercial Anger camera performed at single gamma energy (364 keV) – Han 2008
  Encouraging results:
    Efficiency gain of a factor > 20
    Comparable spatial resolution for the same number of events
OBJECTIVE OF THE PRESENT WORK
Estimate of Compton camera performance in SPECT as a function of:
  Scatterer energy resolution and material
    Main development challenge
  Source activity and random coincidences
Benchmark Anger camera (AC) and Compton camera (CC):
  Comparison of main performance parameters in wide
  energy range (245 keV – 2614 keV)
Simulation study
AC
HOW TO ?

## Slide 16
…image…
AC
Radial distribution from raw data
Background to be studied and handled
…
CC
Coincidence reconstruction
Radial distribution only after reconstruction
…
Radial distributions

## Slide 17
Focus on detector performance in terms of spatial resolution and detection efficiency
Definition of three figures of merit
  Detection efficiency: ratio between reconstructed events and total emitted primary gammas
  Spatial resolution: RMS of radial distribution
  Signal-to-background ratio: ratio between “detected events” and “useful events”
Analysis method
SOURCE:
N primaries
in 4p
g
CC
AC
Transmission + interaction
Interaction + coincidence
M “detected events”
Background
rejection
MLEM reconstruction
Radial distributions: P entries
R
Figures of merit:
1. P/N
2. R
3. P/M
r
#
10 cm

## Slide 18
Data treatment: CC
107 primary gammas simulated in 0.08x4p (camera acceptance cone)
Scatterer-absorber coincidences retrieved with dedicated analysis based on:
  Source activity
  Examination time
  Detector time resolutions
Radial distribution after MLEM reconstruction obtained and analyzed to retrieve the spatial information

## Slide 19
Data treatment: AC
108 primary gammas simulated in 4p
Raw radial distribution analyzed to retrieve the spatial information
Event selection performed by linear fit – background rejection
245 keV
555 keV
1099 keV
1524 keV

## Slide 20: SIMULATION STUDY
Results

## Slide 21
NB: Different scales for the two plots – factor 10
Detection efficiency = reconstructed events / total primaries
Comparison results: detection efficiency
Detection efficiency increased by a factor > 20
for whole investigated energy range
Confirmation of Han results @ 364 keV

## Slide 22
Spatial resolution = RMS radial distribution
Comparison results: spatial resolution
Favorable spatial resolution for energies above 500 keV
Interaction in the patient not considered in this study, reduced at higher energy
Reduction of AC RMS at increasing energy connected to a strong reduction of selected events

## Slide 23
Signal-to-BKG ratio = reconstructed events / detected events
Comparison results: signal-to-BKG ratio
AC optimization for reduced energy range
CC selected events coherent with coincidence study expectations

## Slide 24
Conclusion
Possible application of Compton camera for SPECT further verified
Increased detection efficiency thanks to the “electronic collimation”
  factor > 20 between 245 keV and 2614 keV
Favorable spatial resolution for energies above 500 keV
Wider energy acceptance with the same detector (slight geometrical modifications possible for optimization)
… with Compton cameras
Reduction of injected activity and/or examination time
New high energy sources should be moved to the clinical practice
  Reduced interaction in the patient
  Reduced overall dose
  Increased spatial accuracy
Possibility of control of vectorised therapy
  High energy gamma emission used for imaging
TO DO
Verification tests
  Phantom studies
  Clinical studies
Clinical test of Compton camera prototype
Promising preliminary results
Further study required

## Slide 25
Acknowledgements
CLaRyS Collaboration
L. Balleyguier, E. Bechetoille, D. Bon, A. Bongrand, J.-P. Cachemiche, C. Caplan, L. Caponetto, B. Carlus, X. Chen, M. Dahoumane, D. Dauvergne, D. Delaunay, R. Della Negra, F. Doizon,
N. Freud, M.L. Gallin-Martel, L. Gallin-Martel, N. Giraud, J. Krimmer, D. Lambert, J.M. Létang,
M. Magne, F. Martin, H. Mathez, V. Maxim, G. Montarou, J.-L. Montorio, C. Morel, F. Mounier,
M. Rodo-Bordera, E. Testa, W. Tromeur, Y. Zoccarato
LabEx PRIMES
This work was supported by the LABEX PRIMES (ANR-11-LABX-0063) of Université de
Lyon and by France HADRON (ANR-11-INBS-0007), within the program "Investissements
d'Avenir" (ANR-11-IDEX-0007) operated by the French National Research Agency (ANR).

## Slide 26: Thanks for your attention

## Slide 27: BACKUP

## Slide 28
Radiotracers for comparison study
Decay mode list: e for electron capture, b- for electron emission, b+ for positron emission, IT for isomeric transition. Half-life expressed in days (d), hours (h) or minutes (m). Data extracted using the NNDC On-Line Data Service from the ENSDF database, file revised as of (2017-05-17) [1]
[1] Bhat, M. R., Evaluated Nuclear Structure Data File (ENSDF), Nuclear Data for Science and Technology, page 817, edited by S. M. Qaim (SpringerVerlag), Berlin, Germany, 1992, https://www.nndc.bnl.gov/, 2017-03-27

## Slide 29
ANGER
Radial distributions all energies
COMPTON

## Slide 30
AC data: analysis method validation
Two main tests performed
Check of the number of events in the central collimator holes after fit selection
  Ideal detector
  Comparison with geometrical and physical expectations (gamma absorption in the detector material)
Check the noise-background contribution
  Spatial noise composed by
    Scattered gammas
    Traversing gammas
  Test with infinitely dense collimator – total absorption
    Only gammas traversing the collimator holes are detected
    Qualitative comparison to radial events distribution after selection
Robust analysis method
1099 keV

## Slide 31
Compton camera energy resolution
Scatterer energy resolution
Important parameter for the camera performance
Mainly determined by the Electron Noise Charge (ENC)
2 ENC values tested according to design expectation and first estimate
  ENC = 500 corresponding to about
  = 2 keV
  ENC = 1100 corresponding to about
  = 5 keV
Reduced effect at increasing energy
ENC = 500 chosen for the further study
  Closer to design expectations

## Slide 32
Doppler broadening
Second parameter affecting the energy scatterer resolution
Compton scattering electron approximated to be at rest in reconstruction
The electron energy state affects the energy spectrum – broadening effect
Test of alternative scatterer material for comparison - CdTe
Effect reduced for energies above 1 MeV
Compton camera energy resolution

## Slide 33
Compton camera: random coincidence rate vs source activity
Source activity [MBq]
Effect of random coincidences
Studied as a function of the source activity
Source activity in the range 1-500 MBq – clinical interest
Compatible amount of random and true coincidences at 200 MBq – coherent with clinical routine
200 MBq fixed for further study

## Slide 34
Effect of random coincidences
Random coincidence ratio studied as a function of the source energy
Energies involved in the study
Number of random coincidences decreases at increasing energy, as expected
Random coincidence are partially rejected by the MLEM reconstruction algorithm
Compton camera: random coincidence rate vs source energy

## Slide 35: Particle therapy for cancer treatment (i)
Standard cancer treatment composed by
radiotherapy or/and chemotherapy and/or surgery
Increasing number of patients treated with particle therapy all over the world (p + heavier ions, mostly C)
Many new particle therapy centers built or commissioned
ADVANTAGES (with respect to photons):
Favorable depth-dose distribution – Bragg Peak
  Ballistic targeting of the tumor
  Reduction in the integral dose
    Less dose to healthy tissues
    Reduction of secondary cancer probability and toxicity
Increased Relative Biological Effectiveness (RBE)
  Mostly for ions heavier than protons
    RBEp = 1.1 in clinic
    RBEC = [2-3]
  Possibility to treat radio-resistant tumors (hypoxic)
Durante-Paganetti, 2016, Report of Progress in Physics, Nuclear physics in particle therapy: a review

## Slide 36: Particle therapy for cancer treatment (i.i)
Particle therapy centers in the world,
with focus on Europe
http://www.hirt-japan.info/en/what/japan.html

## Slide 37: Particle therapy for cancer treatment (ii)
Standard clinical routine:
CT scan and treatment planning
  Treatment limits defined to spare the healthy tissues
Fractioned dose delivery
  Two main methods:
    Passive scanning (beam scatters + collimators)
    Active scanning (scanning magnets)
OBJECTIVE
Cover the whole tumor volume
  Bragg peak location determined by the beam energy (Spread-Out Bragg Peak - SOBP)
  Lateral spread to be taken into account
Durante-Paganetti, 2016, Report of Progress in Physics, Nuclear physics in particle therapy: a review

## Slide 38: Particle therapy for cancer treatment (iii)
Standard clinical routine:
CT scan and treatment planning
  Treatment limits defined to spare the healthy tissues
Fractioned dose delivery
  Two main methods:
    Passive scanning (beam scatters + collimators)
    Active scanning (scanning magnets)
OBJECTIVE
Cover the whole tumor volume
  Bragg peak location determined by the beam energy
  Lateral spread to be taken into account
Ballistic features allows a reduction in the irradiation fields
  Further reduction in healthy tissue dose
Durante-Paganetti, 2016, Report of Progress in Physics, Nuclear physics in particle therapy: a review

## Slide 39: Particle therapy for cancer treatment (iv)
DRAWBACKS
Several sources of incertitude
Common to conventional radiotherapy
  Target volume definition
  Intra-fractional changes (tumor shape, size, position, patient weight, positioning, etc.)
  Anatomical uncertainties (lung filling, gas components, bone composition)
Range uncertainties
  CT Hounsfield units conversion to stopping power values -> interest to develop Proton Radiography
  CT imaging uncertainties
  Stopping power measurements / calculation
Biological uncertainties
  RBE definition
NEED FOR ONLINE RANGE MONITORING
Knopf-Lomax, 2013, PMB, In vivo proton range verification: a review

## Slide 40: Ion range monitoring (i)
Online projectile range control needed to fully exploit the dose delivery features of protons and heavier ions (Bragg Peak)
  Reduce treatment planning limitations on volume definition
  Possibly increase the dose per fraction
  Reduce number of fractions and the number of needed irradiation fields
Monitoring possible with secondary radiation produced by nuclear reactions
  High probability:
    ~10 % for protons – 10 cm range
    ~ 40 % for C ions – 10 cm range
  Projectile and target-like fragments (Ex: positron emitters - 11C, 15O)
    PET
  Prompt radiation
    Charged particles -> Interaction Vertex Imaging (IVI)
    Neutrons
    Prompt gammas

## Slide 41: Ion range monitoring (ii)
Prompt radiation emission directly correlated to the dose profile
  Prompt gamma emission distal fall-off correlated to the dose profile and to the Bragg peak position (Min 2006, Testa 2008)
Advantages:
  Emission time < 1 ns after the nuclear interaction – physiological wash-out processes negligible
  Reduced interactions in the patient – millimeter resolution for 108 protons in proton therapy (maximum number of particle per spot)
  (Smeets 2012, Pinto 2014)
Challenge:
  Wide gamma energy range ( ~ 1-10 MeV)
  Large neutron-induced background
GAMMA CAMERA DEVELOPMENT
Dedes 2014
Robert 2013

## Slide 42
Secondary radiation monitoring
Range uncertainties
Paganetti PMB 2012

## Slide 43: PG ion range monitoring: state of the art
Prompt Gamma (PG) monitoring:
  Collimated detection systems
    Knife-edge cameras
    Multi-slit cameras
  Compton cameras – “electronical collimation”
  Non-collimated techniques
    PG timing (PGT)
    PG spectroscopy (PGS)
    PG peak integral (PGPI)
  Several gamma camera projects
  Several prototypes tested in lab
  First prototype in clinical trials (with patients) is by IBA
    Knife-edge camera – Dresden group
      LYSO slabs (4 mm wide, 31.5 mm thick)
      FOV: 10 cm in the Bragg peak region
      Efficiency: 4 x 10-4 (Priegnitz, PMB 2015)
      4 mm fall-off position resolution with 108 protons
PGT
PGPI
PGS
Golnik 2014
Verburg 2015

## Slide 44
Prompt gamma yield
Perali PMB 2014
160 MeV protons in PMMA
Detection with IBA knife edge camera prototype
12C/p prompt gamma yield ratio

## Slide 45: PET ion range monitoring: state of the art
PET monitoring:
  b+ annihilation: 2 back-to-back 511 keV gammas
  Coincidence detection
Several groups working on new prototypes of
In-beam PET
  OpenPET
  DoPET-L
  INSIDE
  TOF PET DPC
  DPGA – CLaRyS group
In beam
Off line
In room
GSI -
HIT -
MGH -
Shakirin 2011

## Slide 46
Compton camera state of the art
Several groups working on Compton camera prototypes
VALENCIA (Llosa, NIMA 2012):
  LaBr3 scatterer and absorber
  ~30 mm absorber size
  Status: In-beam tests
US (Polf 2015):
  CdZnTe scatterer and absorber
  ~40 mm absorber size
  Status: In-beam tests
DRESDEN (Hueso-Gonzalez, JINST 2014):
  CdZnTe scatterer, LSO absorber
  ~30 mm absorber size
  Status: components characterized and tested
MUNICH (Thirolf 2014):
  Si scatterer, LaBr3 absorber
  ~50 mm absorber size
  Status: components characterized and tested

## Slide 47
SCATTERER 7 DSSD (Double-Sided Silicon Detectors) 96x96x2 mm3 planes (10 planes were foreseen at the beginning, 3 are not suitable for the camera)
1.4 mm strip pitch
2 x 64 strips, XY position sensitive
isolated box for thermal regulation
custom electronic cards with new ASIC (developed by IPNL electronics group)
Small prototype tested with radioactive sources
  14 keV FWHM energy resolution @ 122 keV (57Co) at room temperature
  ~20 ns FWHM time resolution P channels
Final detectors characterized in terms of leakage currents
  Tests performed with electronics card first prototype
Detector components: scatterer / collimator

## Slide 48
ABSORBER
LPC Clermont
modular structure
96 BGO streaked blocks (~4 mm spatial resolution)
4 cylindrical PMs (Photo-Multipliers) for each block
ASM front-end cards
BGO blocks from a PET HR+ system by Siemens
  Reconditioning process – scintillating block treatment and PM recoupling
  5 reconditioned blocks available for the first tests
Detector components: absorber
First characterization tests performed on original blocks @ GANIL and with radioactive sources
Energy resolution: 21 % FWHM @ 667 keV (137Cs)
Time resolution: ~2 ns FWHM

## Slide 49
For ion beam therapy monitoring, secondary particle background has to be handled and possibly subtracted
Time Of Flight technique can be employed -> demonstrated possibility to reduce the background level and improve the SNR (Signal to Noise Ratio) – Testa 2010
Fast beam tagging detector needed -> hodoscope under development
SCINTILLATING FIBER HODOSCOPE
128x2 scintillating fibers for 2D information
1 mm2 squared fibers BCF 12 (Saint Gobain) , 140 mm long
2-sided read-out through 1 m optical fibers (one per channel)
8 Hamamatsu PM, 64 channels each (8x8 matrix)
512 channels in total
Optical fibers connection in order to have the two sides of the same fiber on the same PM – uniform response
Custom electronics cards
Remotely controlled moving table as mechanical support, 2 axes
Beam tests 2010-2014
  1 ns FWHM time resolution
  ~90% efficiency
  Rate 107 Hz per PM
TOF and beam hodoscope

## Slide 50
7 custom ASIC cards for scatterer detectors
8 hodoscope DAQ cards – 1 PM each
ASM cards for the absorber
TRIGGER
THOR
EVENT
BUILDER
AMC40
(µ-TCA standard)
Complete acquisition system
Clock

## Slide 51
Mechanical supports
4 movement axes
  3 D overall movement
  2 section distance regulation
Height range from 60 to 150 cm
Wheels for portability
A : thermal box with Silicon detectors / collimator
B : Absorber
C : Positioning table
D : Hodoscope table
2 movement axes
Electronics cards support integrated
D