## Slide 1: Prompt gamma monitoring of proton and carbon therapy: comparative development of a Time-of-Flight Compton camera and a multi-collimated camera
Denis DAUVERGNE
Institut de Physique Nucléaire de Lyon, CNRS/IN2P3 Université Lyon 1
II Symposium on Positron Emission Tomography
Krakow, 20-24 September 2014

## Slide 2
Basic features:
    prompt radiation production by nuclear reactions
    prompt gamma yields and profiles
        Time of Flight issue
Online range verification with prompt gammas
    passive collimation systems
    active collimation: Compton cameras
Outline

## Slide 3: Motivation
Dose deposition during radiotherapy:
  Ionization (uncertainties  margins)
Hadrontherapy:
  Nuclear fragmentation
    High probability
    Influence on dose deposition
    Secondary particles
      g, n, p, fragments
      Radioactive Isotopes (b+)
  Range control by means of
  nuclear reaction products:
      b+ annihilation
      Secondary protons (only carbon)
      Prompt gamma
        ≤ 1 per nuclear reaction
        ~ isotropic emission
        Massive particles: background
GEANT4

## Slide 4
Prompt gamma measurements
with collimated detectors
moving target
beam
collimator
detector

## Slide 5
Prompt gamma measurements
with collimated detectors
moving target
beam
collimator
detector
Timing
monitor
Accelerator HF
or
Synchronization with accelerator HF or monitor: Time of Flight
First evidence of correlation between prompt-gamma profile and ion range: Min et al, APL 2006

## Slide 6: Fast timing and high energy resolution measurements
76mm
J. Verburg, PMB 2013
Fast timing (2ns window)
High spectral resolution (Large volume LaBr + Compton rejection BGO)
Spatial resolution (~1cm FOV)
 10 – 100 counts/ peak for 109 protons
+ Polf PMB 2013:
Oxygen concentration monitoring

## Slide 7: Influence of TOF on PG profiles (collimated cameras)
Roellinghoff PMB 2014
160 MeV protons in PMMA
IBA C230 cyclotron
9.4 ns
310 AMeV carbon ions in PMMA
No TOF
TOF selection
M. Pinto, submitted New J Phys
TOF : mandatory for carbon ions

## Slide 8: Fall-off retrieval precision
Roellinghoff PMB 2014
160 MeV protons in PMMA
IBA C230 cyclotron
Measurement with single (small) detector
Real clinical detector 10 times more efficient  millimetric precision for a distal spot

## Slide 9: Prompt gamma timing
Principle:
Variations of the transient time inside the patient
Measurable with high precision
By means of the prompt-gamma timing profiles
 Measurements at KVI:
Clinical applicability: 2mm range deviation measurable within few seconds
C. GOLNIK, PMB 2014

## Slide 10
What do we want ?
Online control with prompt gammas

## Slide 11
What do we want ?
- Range verification with mm accuracy
    For single pencil beam spot (distal)
      Protons: 108 particles
      Carbon ions: 106 particles
Online control with prompt gammas

## Slide 12
What do we want ?
- Range verification with mm accuracy
    For single pencil beam spot (distal)
      Protons: 108 particles
      Carbon ions: 106 particles
    For distal energy slide (statistics x10)
Online control with prompt gammas

## Slide 13
What do we want ?
- Range verification with mm accuracy
    For single pencil beam spot (distal)
      Protons: 108 particles
      Carbon ions: 106 particles
    For distal energy slide (statistics x10)
    For whole fraction (Statistics x1000) or passive delivery
Online control with prompt gammas

## Slide 14
What do we want ?
- Range verification with mm accuracy
    For single pencil beam spot (distal)
      Protons: 108 particles
      Carbon ions: 106 particles
    For distal energy slide (statistics x10)
    For whole fraction (Statistics x1000) or passive delivery
- real time?
- 2D or 3D spatial information?
- target composition: spectral information
PG yield above 1 MeV : ~ 0.3% /cm per proton, ~ 2% /cm per carbon
 Compromise with statistics
Online control with prompt gammas

## Slide 15: Main features - Challenges
Relatively large number of PG emitted ( >108 per fraction)
Correlated to ion range
Real time information
Poly-energetic
E> 1MeV : minimum absorption
  Escape from patient
  Difficult to collimate and detect
  Current SPECT devices not adapted:
  New technologies/concepts needed
Large background (neutrons…)

## Slide 16: Collimated cameras
2 kinds of cameras
  Multi-slit (Lyon, Delft, Seoul, IBA)
  Knife-edge (Seoul, Delft, IBA)
+ Simplicity
+ Large field of view for multi-collimated
- Mechanical collimation…
IBA Knife edge camera :
under clinical tests: millimetric precision
at pencil beam scale.

## Slide 17: Multislit TOF-gamma camera

## Slide 18: Compton cameras
Group |  |  | stages | scatterer | absorber | Efficiency (10-5) | Resolution (mm) | Proto/simu | ref
Lyon |  |  | 3 | DSSD | LYSO | ~1 | 6 | simu | Richard 2010
Baltimore -Texas |  |  |  | Ge CdZnTe | Ge/LaBr CdZnTe | ~1 ? | ? | Simu proto | Peterson 2010
Seoul |  |  |  | DSSD | NaI | ? | ~12 | Proto tested | Seo 2011
ENVISION | Valencia |  | 2 -3 | LaBr | LaBr | ? | 7.8 | Proto tested | Llosa 2013
 | Lyon |  | 2 | DSSD | BGO | ~25 | ~8 | Simu/proto in prog | Roellinghoff 2011, Ley 2014
 | Dresden |  |  | CdZnTe | LSO/BGO | ? |  | Proto tested | Hueso- Gonzalez 2014
Munich |  |  |  | DSSD | LaBr | ~1-100 | ~8 | Proto in prog | Thirolf 2014
ETCC |  | Kyoto | Gas: Ar+C2H6 |  | GSO | 0.3 | ? | Proto tested | Kurosawa 2012
 |  | Seoul | g conv.+hodo |  | CsI | 0.6 | 100 | Simus | Kim 2012
No collimation: potentially higher efficiency
Potentially better spatial resolution (< 1cm PSF)
If beam position known  simplified reconstruction
3D-potential imaging (several cameras)

## Slide 19
Jean-Luc LEY, PTCOG 2014

## Slide 20: Silicon double-sided strip detectors
J. Krimmer, ANNIMA 2014

## Slide 21: DSSD front-end electronics

## Slide 22: BGO absorber

## Slide 23: Precise timing: fast monitor Scintillating fibers hodoscope
1mm2 square fibers (128 +128 fibers)
Light transmission: optic fibers
Photomultiplier : Hamamatsu H8500
Home-made ASICs electronic readout : discriminator + TDC at 108 Hz rate capability (S. Deng 2011)
Tests:
0.5 ns resolution
Admissible dose > 1012 carbon ions/cm2
Count rate < 4x106 Hz per PMT: use of several PMTs
Prototype

## Slide 24: Construction of full-size prototype (collaboration IPN Lyon, CREATIS Lyon, LPC-Clermont, CPPM-Marseille)

## Slide 25: Compton-camera count rate issue
Simulation: line-cone reconstruction for Lyon prototype
1 distal spot (108 incident protons) incident on PMMA target, 160 MeV
Pulsed beam (IBA C230)
Clinical intensity: Reduced intensity:
200 protons/bunch 1 proton/bunch
Absorber count rate: 200MHz
JL Ley, PTCOG 2014

## Slide 26: Concluding remarks
Prompt gamma :
emerging technique close to clinical translation
The most appropriate modality for real time control of the range during proton treatment
  Collaborative effort during ENVISION (FP7), worldwide competition
Collimated systems
  Millimetric range-control at the pencil-beam scale for protons
  First prototype tested in clinical conditions (knife-edge IBA)
  Multi-collimated cameras: similar performances, better potentialities
Compton cameras: still under development
  Lyon prototype with TOF: comparable statistics with collimated devices, spatial resolution, 2-3D imaging
  Necessity to work at reduced intensity
  New concepts (with calibration issue): PG timing, PG spectroscopy
Accelerator-dependent devices (count rates, TOF)

## Slide 27: Special thanks
Collaborators:
CPPM Marseille:C. Abellan, J.P. Cachemiche, C. Morel
IPN Lyon: L. Caponetto, X. Chen, M. Dahoumane, J. Krimmer, J.-L. Ley, H. Mathez, M. Pinto, E. Testa, Y. Zoccarato
CREATIS Lyon: N. Freud, J.M. Létang,
LPC-Clermont: D. Lambert, M. Magne, G. Montarou
Acknowledgements:
France Hadron, FP7 ENTERVISION, FP7 ENVISION, FP7 ULICE

## Slide 28: Compton cameras
Group |  |  | stages | scatterer | absorber | Efficiency (10-5) | Resolution (mm) | Proto/simu | ref
Lyon |  |  | 3 | DSSD | LYSO | ~1 | 6 | simu | Richard 2010
Baltimore -Texas |  |  |  | Ge CdZnTe | Ge/LaBr CdZnTe | ~1 ? | ? | Simu proto | Peterson 2010
Seoul |  |  |  | DSSD | NaI | ? | ~12 | Proto tested | Seo 2011
ENVISION | Valencia |  | 2 -3 | LaBr | LaBr | ? | 7.8 | Proto tested | Llosa 2013
 | Lyon |  | 2 | DSSD | BGO | ~25 | ~8 | Simu/proto in prog | Roellinghoff 2011, Ley 2014
 | Dresden |  |  | CdZnTe | LSO/BGO | ? |  | Proto tested | Hueso- Gonzalez 2014
Munich |  |  |  | DSSD | LaBr | ~1-100 | ~8 | Proto in prog | Thirolf 2014
ETCC |  | Kyoto | Gas: Ar+C2H6 |  | GSO | 0.3 | ? | Proto tested | Kurosawa 2012
 |  | Seoul | g conv.+hodo |  | CsI | 0.6 | 100 | Simus | Kim 2012
Lyon project:
TOF and beam position with hodoscope
Large size camera

## Slide 29: Prompt gamma measurements
PG yield above 1 MeV : ~ 0.3% /cm per proton ~ 2% /cm per carbon
110 MeV protons in water
95, 300, 310 MeV/u carbon in H20 and PMMA
M.Pinto, submitted New J Phys
J.Verburg, PMB 2013
E> 1MeV

## Slide 34: Prompt gamma yields: heterogeneous targets
High resolution profiles: influence of heterogeneities close to the Bragg peak
M. Pinto et al, submitted to Med Phys
95 MeV/u carbon ions

## Slide 35
Prompt gamma measurements
with collimated detectors
Energy: <1 MeV to 10 MeV
A small fraction is measured as discrete lines
Low energy gammas: larger scattered fraction
Energy spectrum 160 MeV protons in PMMA, NaI(Tl) detector
Smeets PMB 2012
moving target
beam
collimator
detector